## Slide 1: TMS–EMG toolbox, user manual
Lari Koponen
2021-04-21

## Slide 2: Overview
TMS–EMG toolbox connects
  TMS device (MagVenture MagPro)
  EMG amplifier (BrainVision ExG)
  On-line TMS–EMG analysis
BNC cable
2N+1 electrodes
PARTICIPANT
Neuronavigation
(BrainSight)
Serial cable*
COM_**
COM2
Proprietary connector
EMG
(ExG)
TMS
(MagPro)
TrigOUT
TrigIN
TrigIN
* Or, a USB-to-serial adapter and a serial cable, works too
** For the adapter on current computer, COM3; for the port on the motherboard COM1
PC, with:
Brain Vision Recorder (software)
TMS–EMG toolbox (MATLAB app)
(optional)
Trigger IN box
Footswitch
USB cable
(optional) BNC cable

## Slide 3: Usage 1/5
Recommended use pattern
  Turn on TMS device
  Turn on ExG
  Start BrainVision recorder
  SEND TEST PULSES
    BrainVision Recorder sometimes does not capture the first trigger after turn-on, we do not want this to occur during recording

## Slide 4: Usage 2/5
Recommended use pattern
  Start MATLAB
    Add ‘TMS-EMG toolbox 20210419’ to path
    Navigate to a folder where you want to record the data
    Run command ‘TMSEMG’

## Slide 5: Usage 3/5
Recommended use pattern
  On the ‘Connect ExG’ tab, press connect
    This takes 15–20 s
    In case an error popup shows, about ‘Parallel Processing Toolbox’, and you are not running a second copy of the software, push the red button
  On the ‘Connect TMS’ tab, name the (raw) log file for TMS pulses and press connect
    In case an error popup shows, serial port is already in use (run ‘instrreset’, check device manager)

## Slide 6: Usage 4/5
Recommended use pattern
  Use ‘Acquire IO curve’ tab
    One pulse per intensity between min & max
    Two pulses between dense min & dense max
    Random seed for pseudorandom sampling sequence (pulse are saved into the ‘EMG’ data output)
    Channel number (just for note, saved to ‘EMG’ data output)
    ISI, uniformly random between limits
    Start: start sampling
    Pause sampling (resume with start)
    Stop: cancel sampling (permanently, start starts from beginning)
IO curve acquisition can be started or paused with a press of the footswitch.
  The footswitch connects to the OPTIONAL trigger input box, which further (optionally) takes trigger output signal from the TMS device.
  If this signal is present, the output data file will contain field ‘triggerIN’ with a ‘datetime’ object for each trigger signal by the TMS device, this data can be used to resolve the ambiguity with paired pulses
    The on-line trigger counts follow the serial data, the correction is left to the user off-line, POST session
  These signals are generated even if the TMS device is not connected in software, as long as the TMSEMG window is active!

## Slide 7: Usage 5/5
Saving data
  ‘Save EMG data’ button
    Saves all data acquired during current session to ‘.mat’ file
  If not saved, asks on close
Example of (synthetic) error
  COM3 is already in use
    By running the TMS device backend manually
    Solvable by ‘delete()’ of that object

## Slide 8: Output data formats
EMG data file (.mat)
  ‘description’ the date the data was saved
  ‘data’, struct with cell arrays of
    ‘TMS’ for each TMS pulse
      ‘clock’ (datetime object)
      ‘amplitude’ (% MSO)
      ‘didt’ (MA/s)
      ‘mode’ (0=standard, 1=power, 2=twin, 3=dual)
      ‘waveform’ (0=monophasic, 1=biphasic, 2=halfsine, 3=biphasic burst)
    ‘EMG’ for each acquired epoch
      {‘clock’ ‘raw data [8×2751]’ ‘filtered data [8 8×2751]’ ‘classification [8×2]’}
        Data, rows of FILTERED EMG data from –0.2 to 0.2 s
        Classification rows of MEP amplitude and MEP classification (1=MEP, 0=no MEP, –1=rejected trial)
    ‘IO’ for each started IO curve
      {‘clock’ ‘pulse sequence [1×N]’ ‘TMS index at start’ ‘EMG index at start’}
TMS device logfile (.txt)
  Plain text of all data outputted by the TMS device
  Event types 1-3, described in the MagPro user manual, have been decoded, the rest are just saved
  Mostly for FYI, and for records
  Contains every amplitude adjustment made on the TMS device front panel etc.
  Data key provided in the class help of ‘MagProController’

## Slide 9: Known limitations
Neuronavigation is not coupled to the on-line loop
  Cannot do “Nexstim-like” features, like ‘trigger-only-if-coil-is-within-tolerance-of-the-correct-target’
Serial data from MagPro has undefined behavior
  Dual/twin mode sends “randomly” one or two pulse messages
    Case point: PP with 2.9 ms IPI sends RANDOMLY one or two messages
  Solutions:
    Extend BNC to PC, the triggers are deterministic
      Via a parallel port (or via Arduino combining the two data streams into one…)
    Make experiment specific assumptions of pulse parameters, and try to deconstruct what pulses have actually happened